HCD6817ZCompact Weather Station

（T+RH+P+WD+WDIR+Rianfall+Illuminance）

OVERVIE

HCD6817Z is designed and produced by Handan Yunnong Intelligent Agricultural Technology Co., Ltd. It innovatively realizes the meteorological standard parameters （T、RH、P、WD、WDIR、Rianfall、Illuminance）through a highly integrated structure, which can realize the 24-hour continuous online monitoring of outdoor meteorological parameters, and output the 7 parameters to the user at one time through the digital communication interface.

APPLICATION

HCD6817Zcan be widely used in urban environmental monitoring, wind power generation, meteorological monitoring, Bridges, tunnels, ships, aviation airports, agricultural meteorology, water conservancy meteorology, power environment, highway meteorological monitoring and other fields.No maintenance and on-site calibration is required.

FEATURES

1、Standard configuration monitoring T、RH、P、WD、WDIR、Rianfall、Illuminance, RS485（MODBUS ）；

2、High precision, reliable performance, suitable for outdoor weather harsh environment;

3、Parameter collection, optional wireless data collector, automatic data upload network platform, mobile phone client real-time data view；

4、Real-time monitoring of meteorological environmental data, low cost, suitable for grid distribution；

5、Small size, modular design, flexible layout；

6、Data acquisition adopts 32-bit high-speed processing chip, stable and anti-interference.

TECHNICAL SPECIFICATION

| Measture item | Principle | Range | Resolution | Accuracy |
| --- | --- | --- | --- | --- |
| Atmospheric temperature | Energy gap method | -40-60℃ | 0.1℃ | ±0.3℃（25℃） |
| Atmospheric humidity | Capacitance method | 0-100%RH | 0.1% | ±3%RH |
| Atmospheric pressure | Piezoresistive | 300-1100hpa | 0.1hpa | ±0.5hpa（0-30℃） |
| Wind speed | Ultrasonic | 0-60m/s | 0.01m/s | （（0-30m/s）±0.3m/s or±3% |
| Wind direction | Ultrasonic | 0-360° | 0.1° | ±2° |
| Rianfall | Photology | 0-200mm/h | 0.2mm | Error<10% |
| Illuminance | Photodiode | 0-100KLUX | 10LUX | 3% |
| Supply | 12-24VDC, solar power |  |  |  |
| Output | RS485(Modbus) |  |  |  |
| fixed form | - Sleeve fixed； B)Flange adapter fixed |  |  |  |
| Power consumption | ＜1W@12V |  |  |  |
| Ingress Protection | IP66 |  |  |  |
| Main material | ASA+engineering plastics |  |  |  |
| Mounting frame | 1.5m, 3m Flange, 1.8m tripod |  |  |  |
| GPS | Device tracking and positioning |  |  |  |
| Data Release | PC、Telephone APP、LED、LCD TV display, etc |  |  |  |

1.Rainfall sensor (optics)

2.illuminance/solar radiation

3.Arrows to point north

4、Ultrasonic probe

5、Control circuits

6、Thermometer shelter (temperature, humidity, air pressure monitoring position)

7、Bottom fixed with flange

※This product can be equipped with electronic compass, GPRS(built-in)/GPS (optional)

MODBUS RTU COMMUNICATION PROTOCOL

Baud rate：9600

Data bits：8

Stop bit：1

Check bit：None

5.1 CRC instruction：

Among all the following instructions, the two bytes of CRC16 in MODBUS RTU protocol are as follows: the low byte comes before the high byte comes after.

In the following instructions, the sensor address 0xFF is assumed (the sensor default address is FF)

5.2 Return error code rule：

When receive error instruction (including CRC16 validation error), the sensor will not return error code. After the instruction is issued 100ms, it fails to receive issuing instruction The upper computer may consider failure and may resend the instruction.

5.3 Standard MODBUS register specification

Note: The quantity or length of register in Modbus is two bytes and 16 bits(the high byte comes first and the low byte comes last), Non - byte 8 bits are a unit.

The user shall ensure that the range of the address and the range of quantity of registers are within  the system range.If it exceeded, the output of the sensor will be unpredictable. The user shall ensure that the MODBUS  meets the requirements of this manual in the software design of the upper computer

Input register: read with function code 03

| Address | operation | Contents | Notes |
| --- | --- | --- | --- |
| 0x0001 | read-only | Noise, Hexadecimal number is magnified as 10 times, such as 0x021C indicates noise of 54.0db |  |
| 0x0002 | read-only | Keep |  |
| 0x0003 | read-only | SO2 concentration,hexadecimal ,such as 0x0172, indicates that SO2 concentration is 370ppb. |  |
| 0x0004 | read-only | NO2 concentration,hexadecimal ,such as 0x0036, indicates that NO2 concentration is 54ppb |  |
| 0x0005 | read-only | CO concentration,hexadecimal ,such as 0x0A00, indicates that CO concentration is 2560ppb |  |
| 0x0006 | read-only | O3 concentration,hexadecimal ,such as 0x0123, indicates that O3 concentration is 2560ppb |  |
| 0x0007 | read-only | PM2.5 concentration,hexadecimal ,such as 0x0172, indicates that PM2.5 concentration is 370ug/m³ |  |
| 0x0008 | read-only | PM10 concentration,hexadecimal ,such as 0x0193, indicates that PM10 concentration is 403ug/m³ |  |
| 0x0009 | read-only | Atmospheric temperature, Hexadecimal number plus 40 then it is magnified as 100 times, such as 0x1B00 indicates 6912/100-40=29.12℃ |  |
| 0x000A | read-only | Atmospheric humidity, Hexadecimal number is magnified as 100 times, such as 0x1603 indicates 5635/100=56.35% |  |
| 0x000B | read-only | Atmospheric tempress, Hexadecimal number is magnified as 10 times, such as 0x2784 indicates 10116/10=1011.6hPa |  |
| 0x000C | read-only | Wind speed, Hexadecimal number is magnified as 100 times, such as 0x0125 indicates 293/100=2.93m/s |  |
| 0x000D | read-only | Wind direction, Hexadecimal number is magnified as 10 times, such as 0x0C14 indicates 3092/10=309.2° | North is 0 ° |
| 0x000E | read-only | Rainfall, Hexadecimal number is magnified as 10 times, such as 0x0C14 indicates 2.2mm | The default is 10 minutes of rain |
| 0x000F | read-only | solar radiation, Hexadecimal number, such as 0x0172 indicates 370W/㎡ |  |
| 0x0010 | read-only | illumination, Hexadecimal number is magnified as 100 times, such as 0x0123 indicates 2.91.Klux |  |
| 0x0011 | read-only | ultraviolet ray, Hexadecimal number, such as 0x05 indicates 5 |  |
| 0x0012 | read-only | CO2, Hexadecimal number, such as 0x01F4 indicates 500ppm |  |
| 0x0013-0x001f | read-only | Keep |  |
| 0x0105 | read-only | electronic compass, Hexadecimal number, such as 0x0036 indicates 54° | North is 0 |

Internal register: function code 03 to read; function code 06 to write.

| Address | operation | Contents | Notes |
| --- | --- | --- | --- |
| 0x0000 | Read and write | Sensor address, range 0x01-0xff (decimal), factory set to 0xFF |  |

Note：The data of address indicated by hexadecimal, for example, address 20 (decimal) read and write values is 0x14

5.4 communication examples

Here is an example of using Modbus RTU command to access system registers

1.Read multiple input registers (5 live data) commands

Send：FF 03 00 09 00 05 40 15

| FF | 03 | 00 09 | 00 05 | 40 15 |
| --- | --- | --- | --- | --- |
| System address | function code | Register address | Register number | CRC16 check bit generated automatically by software |

Answer：FF 03 0A 1A 57 0C 5F 27 83 00 00 0C 14 36 E0

| FF | 03 | 0A | 1A 57 0C 5F 27 83 00 00 0C 14 | 36 E0 |
| --- | --- | --- | --- | --- |
| System address | function code | Number of bytes of data segment | Data segment data | CRC16 check bit |

parse data：

0x1A57= 0x1A * 256 + 0x57= 6743

Temperature = 6743/100-40=27.43℃

0x0C5F= 0x0C * 256 + 0x5F= 3167

Humidity=3167/100=31.67%RH

0x2783= 0x27 * 256 + 0x83= 10115

atmospheric pressure=10115/10=1011.5hPa

0x0000=0

Wind speed=0/100=0.00m/s

0x0C14=0x0C * 256 + 0x14= 3092

Wind direction=3092/10=309.2°

- 2.Read the single input register command

- Send：FF 03 00 01 00 01 C0 14

| FF | 03 | 00 01 | 00 01 | C0 14 |
| --- | --- | --- | --- | --- |
| System address | function code | Register address | Register address | CRC16 check bit generated automatically by software |

Answer：FF 03 02 02 1C 91 39

| FF | 03 | 02 | 02 1C | 91 39 |
| --- | --- | --- | --- | --- |
| System address | function code | Number of bytes of data segment | Data segment data | CRC16 check bit |

parse data：

0x021C= 0x02 * 256 + 0x1C= 540

Noise = 540/10=54.0dB

- 3.Read the address register command

Send：00 03 00 00 00 01 85 DB

| 00 | 03 | 00 00 | 00 01 | 85 DB |
| --- | --- | --- | --- | --- |
|  | function code | Register address | Register address | CRC16 check bit generated automatically by software |

Answer：00 03 02 00 01 44 44

| 00 | 03 | 02 | 00 01 | 44 44 |
| --- | --- | --- | --- | --- |
|  | function code | Number of bytes of data segment | Data segment data | CRC16 check bit |

Data segment data0x0001 =01   system address 01

4、Modify internal register（System address）Command（Modify address to 0x33）

Send：00 06 00 00 00 33 C8 0E

| 00 | 06 | 00 00 | 00 33 | C8 0E |
| --- | --- | --- | --- | --- |
|  | function code | Register address | New address | CRC16 check bit |

Answer：00 06 00 00 00 33 C8 0E  （Indicates successful modification）

| 00 | 06 | 00 00 | 00 33 | C8 0E |
| --- | --- | --- | --- | --- |
|  | function code | Start address | New address | CRC16 check bit |

5、 Read the electronic compass (output value is the Angle between compass north and sensor north arrow)

Send：FF 03 01 05 00 01 80 29

| FF | 03 | 01 05 | 00 01 | 80 29 |
| --- | --- | --- | --- | --- |
| Station number | function code | Register address | Register number | CRC16 check bit |

Answer：FF 03 02 00 36 11 86

| FF | 03 | 02 | 00 36 | 11 86 |
| --- | --- | --- | --- | --- |
| Station number | function code | Number of bytes of data segment | Data segment data | CRC16 check bit |

6、Set time of rainfall accumulation

Send：00 06 01 04  00 0A  48 21

| 00 | 06 | 01 04 | 00 0A | 48 21 |
| --- | --- | --- | --- | --- |
|  | function code | Start address | Time accumulation(min) | CRC16 check bit |

Answer：00 06 01 04  00 0A  48 21 （Indicates successful modification)

| 00 | 06 | 01 04 | 00 0A | 48 21 |
| --- | --- | --- | --- | --- |
|  | function code | Start address | accumulation(min | CRC16 check bit |

Note：The default immediate time is 10 mins from factury

When the cumulative time of rainfall is set to 0, the equipment does not automatically reset , but accumulates all the time. The rainfall can be reset by resetting the cumulative time of rainfall after power failure

After done sets restart the prodcuts

- 7、Set the declination correction（Only for the selection of electronic compass）

Send：00 06 01 03  00 05  B9 E4

| 00 | 06 | 01 04 | 00 05 | B9 E4 |
| --- | --- | --- | --- | --- |
|  | function code | Register address | correction angle | CRC16 check bit |

Answer：00 06 01 03  00 05   B9 E4  （Indicates successful modification）

| 00 | 06 | 01 03 | 00 05 | B9 E4 |
| --- | --- | --- | --- | --- |
|  | function code | Start address | correction angle | CRC16 check bit |

- Correction Angle

- The high octet represents the correction direction, 0x00 represents a positive correction, and 0x01 represents a negative correction

- The lower eight are the angles that need to be corrected

- Such as 0x00 05 Indicate to need to increase the output by 5 degrees

- 0x01 03 Indicates to need to reduce the output by 3 degrees

- 8、GPRSSet up network information (for the selection of GPRS)

Set up IP   AT+GPRSIP=xxx.xxx.xxx.xxx#

Setport AT+GPRSPORT=xxxxx#

Save and effect the Settings   AT+GPRSRESET

Read IP、Port     RD+GPRSADD=

- 9、Read the latitude and longitude information

- Send AT+GPS#

- Answer GPS:36.12345,N;114.12345,E#

- 5.5 Else

customized according to user requirement, specific requirements can consult the relevant sales.